## Sheet1
Problem | Description | Code Location (Class, Function name etc.) | Code Snippet | Output to user | Testing Procedure | Last tested
Network Connection | Ensures that prior to each perturbation, a connection between the perturbation software and the treadmill remote control is still active. | PerturbationTabWidget::setTryNotToStep:1198-1222 | if(checkSocketConnection() …) | An alert dialog box will be displayed when the user tries to run a perturbation | (i) Start pert software with remote control software on; establish connection between pert software and remote control (i.e., press connect); run perturbation; turn off remote control; run perturbation; repeat 10 times. Error should be issued with disconnection and perturbation should not occur. (ii) (a) Start pert software without remote connection on; try to connect; an error should occur. (b) Re-establish remote connection and run pert. Software should run without issue.
Mouse Connection | Checks if mouse is connected before running perturbation | PerturbationTabWidget::setTryNotToStep:1198-1222 | If(… && mouseInterface->GetDevices()) | An alert dialog box will be displayed when the user tries to run a perturbation
Mouse Response | if treadmill movement signal is not received from mouse, after 1.5 seconds, a deceleration signal is sent to the treadmill and a warning window is displayed | PerturbationTabWidget::setTryNotToStep:766,960-972,1919-1923 | QTimer::singleShot(safetyTimeout, this, SLOT(safetyStop()));void PerturbationTabWidget::safetyStop() { if(!mouseInterface->movementDetectedBool && !slotTimeOutRunBool) { startDecelTimer(); QMessageBox::StandardButton criticalErr = QMessageBox::critical( this, APP_NAME, tr("Treadmill movement was not detected by the tachometer; perturbation has been stopped.") ); slotTimeOutRunBool = false; } }Void PerturbationTabWidget::setSafetyTimeout(double time) { safetyTimeout = time; qDebug() << safetyTimeout; } | A warning dialog is issued to the operator that the mouse did not detect movement | (i) With mouse connected, run 50 perturbations to identify any potential software flaws. (ii) Disconnect mouse from USB after pert software is turned on; try to run a perturbation. Pert should run but stopped by failsafe.
Accelerate/Decceleration Tables Synced | Check is performed to make sure that the data in the perturbation table, database, and acceleration/deceleration boxes are in synchrony | PerturbationTabWidget::setTryNotToStep:1196-1216 | if(prp->compareDbWithRunTable(returnTypeNum(runTableWidget->item(currentRunRowIndex, typeCol)->text()), returnTypeNum(runTableWidget->item(currentRunRowIndex, levelCol)->text()))) { mouseInterface->WriteEvent(); //Starts Cameras int tryNotToStepTimeout = 2000; QString tryNotToStep("TRY NOT TO STEP"); sInterface->updateTextField(tryNotToStep); sInterface->updateTextColor("white"); QTimer::singleShot(tryNotToStepTimeout, this, SLOT(randomDelay())); } else { QMessageBox::StandardButton criticalErr = QMessageBox::critical( this, APP_NAME, tr("The Acceleration Time in the Run Table is out of Sync with the Database.\nPerturbation has been Terminated.") ); } | An alert dialog box will be displayed when the user tries to run a perturbation
Deceleration | After deceleration is complete, another deceleration signal is sent to make sure that the treadmill has stopped | void PerturbationTabWidget::slotTimeout():919 | QTimer::singleShot(repoFeetTimeout, this, SLOT(setPositionFeet())); | No output, logs indicate commands being sent | (i) Run 10 perturbations and determine command is being send to treadmill via console interface.
Prevention additional treadmill commands | During perturbations, buttons are greyed out to prevent users from sending additional acceleration signals while a perturbation is being run. Several flags are used to prevent multiple acceleration signals from being sent. | PerturbationTabWidget::randomDelay():715-729 | startPertRunBtn->setDisabled(true); nextPertRunBtn->setDisabled(true); prevPertRunBtn->setDisabled(true); | Buttons are greyed out
Error feedback | All safety critical functions are wrapped in a try/catch block to alert the user of problems |  |  | An alert dialog box will be displayed when the user tries to run a perturbation
| Mouse will trigger in response to inadvertent movement on the treadmill |  |  | User won’t notice anything